**Sapelo Island** (SAP) **NERR Meteorological Metadata**

**January-December 2018**

**Latest Update:** 10/28/2019

**I. Data Set and Research Descriptors**

1. **Principal investigator(s) and contact persons –**

[**1) Principal investigator(s) and contact persons** Rachel Guy, Research Coordinator

E-mail: rachel.guy@dnr.ga.gov (912) 485-2251

Patrick Hagan, Research Assistant, bluewatermotion@hotmail.com

**Address:**

State of Georgia Natural Resources Dept.

1 Long Tabby Lane

Sapelo Island, GA 31327

(912)-485-2251

**2) Entry verification –**

Data are uploaded from the CR1000 data logger to a Personal Computer (IBM compatible). Files are exported from or LoggerNet in a comma-delimited format and uploaded to the CDMO where they undergo automated primary QAQC and become part of the CDMO’s online provisional database. During primary QAQC, data are flagged if they are missing or out of sensor range. The edited file is then returned to the Reserve where it is opened in Microsoft Excel and processed using the CDMO’s NERRQAQC Excel macro. The macro inserts station codes, creates metadata worksheets for flagged data and summary statistics, and graphs the data for review. It allows the user to apply QAQC flags and codes to the data, append files, and export the resulting data file to the CDMO for tertiary QAQC and assimilation into the CDMO’s authoritative online database. For more information on QAQC flags and QAQC codes, see Sections 11 and 12. Patrick Hagan is responsible for all data management.

1. **Research objectives –**

The principal objectives are to record meteorological information for the Sapelo Island NERR’s site that can be used 1) as a reference for meteorological data for research projects on the Reserve, 2) to give meteorological context for our half hourly SWMP water quality data, and other long term environmental monitoring projects at the Reserve, 3) to observe and characterize important events such as storms, heat and cold waves, droughts, and heavy rainfalls, and 4) to detect trends and characterize climate variability over the long-term

**4) Research methods –**

Campbell Scientific data telemetry equipment was installed at the Marsh Landing station on 02/15/2007 and transmits data to the NOAA GOES satellite, NESDIS ID #3B036592. The transmissions are scheduled hourly and contain four (4) data sets reflecting fifteen-minute data sampling intervals. Upon receipt by the CDMO, the data undergoes the same automated primary QAQC process detailed in Section 2 above. The “real-time” telemetry data become part of the provisional dataset until undergoing secondary and tertiary QAQC and assimilation in the CDMO’s authoritative online database. Provisional and authoritative data are available at [http://cdmo.baruch.sc.edu](http://cdmo.baruch.sc.edu/).

Data are reported in Eastern Standard Time(EST) for the year.

The 15 minute Data are collected in the following formats for the **CR1000**:

Averages from 5-second data:

Air Temperature (°C), Relative Humidity (%), Barometric Pressure (mb), Wind Speed (m/s), Wind Direction (degrees), Battery Voltage (volts)

Maximum and Minimum Air Temperature (°C) and their times from 5-second data (these data are available from the Reserve)

Maximum Wind Speed (m/s) and time from 5-second data

Wind Direction Standard Deviation (degrees)

Totals:

Precipitation (mm), PAR (millimoles/m2), and Cumulative Precipitation (mm) (Cumulative precipitation is no longer available via export from the CDMO. Please contact the Reserve or the CDMO for more information or to obtain these data.)

Sensors are visually inspected on a monthly basis and are removed and recalibrated by the manufacturer on the following schedule:

- Temperature/Humidity- yearly recalibration

- Rain Gauge- yearly recalibration

- Wind Speed/Direction- yearly or every 2 years (depending on the sensor)

- Barometric Pressure- every 2 years recalibration

- PAR- every 2 years recalibration

- CR1000-every 5 years (required beginning 2014, one year initial grace period)

**5) Site location and character –**

The site is located at 31° 25' 4.08 N, 81° 17' 43.26 W, about 15 feet (4.6m) above sea level. All sensors are mounted on a 10m-aluminum tower located in the northern corner of the Marsh Landing parking lot on the southwest corner of the island. The parking lot surface is approximately a meter above the surrounding marsh. The heights of the sensors on the tower are as follows:

Temperature and relative humidity 2m

Barometric pressure 1m

PAR 3m

Precipitation gauge 6m

Wind sensor 10m

It is bordered with salt marsh to the north and south with transition into pine forest occurring 1/2 mi to the east. On the immediate west lie the Duplin River and the ferry dock, which is also the location of our lower Duplin and marsh Landing Water Quality sampling sites (about 202ft (61.6m) from the MET station). The station is well exposed to all winds and weather with little blockage and no shading. This region is subject to multiple severe weather phenomena partially due to the proximity to the ocean. These phenomena include severe summer thunderstorms, which can cause drastic, localized drops in pressure, temperature, and heavy rains; powerful fall and winter frontal systems carrying prolonged strong winds (usually NNE), drastic and sudden drops in temperature and pressure, and long, steady rains; and finally the late summer and early fall hurricanes. It must be noted that due to the remote location and proximity to the ocean our weather patterns can vary greatly from those on the mainland, particularly temperatures which tend to moderate due to our being surrounded by water. The nearest sites for comparison is the Grays Reef NOAA weather Buoy located about 20nm east of Sapelo Island and Glynco airport located approx. 30 mi to the SSE.

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Station Code | SWMP Status | Station Name | Location | Active Dates | Reason Decommissioned | Notes |
| SAPMLMET | P | Marsh Landing | 31° 25' 4.08 N, 81° 17' 43.26 W | 09/01/2002 - | NA | NA |

**6) Data collection period –**

January 1st at 00:00 to December 31st at 23:45 2018

|  |  |
| --- | --- |
| Raw File Start Date and Time | Raw File End Date and Time |
| 12/13/2017 13:00 | 01/24/2018 13:00 |
| 01/24/2018 13:15 | 04/19/2018 13:30 |
| 04/01/2018 00:00 | 04/19/2018 13:30 |
| 04/19/2018 13:35 | 05/08/2018 12:15 |
| 05/08/2018 12:30 | 05/10/2018 11:30 |
| 05/10/2018 00:00\* | 06/30/2018 23:45 |
| 07/01/2018 00:15 | 09/30/2018 23:45 |
| 10/01/2018 00:00 | 12/31/2018 23:45 |

\*Telemetry data used to create this raw file

**7) Distribution –**

NOAA retains the right to analyze, synthesize and publish summaries of the NERRS System-wide Monitoring Program data.  The NERRS retains the right to be fully credited for having collected and process the data.  Following academic courtesy standards, the NERR site where the data were collected should be contacted and fully acknowledged in any subsequent publications in which any part of the data are used.  The data set enclosed within this package/transmission is only as good as the quality assurance and quality control procedures outlined by the enclosed metadata reporting statement.  The user bears all responsibility for its subsequent use/misuse in any further analyses or comparisons.  The Federal government does not assume liability to the Recipient or third persons, nor will the Federal government reimburse or indemnify the Recipient for its liability due to any losses resulting in any way from the use of this data.

Requested citation format:

NOAA National Estuarine Research Reserve System (NERRS). System-wide Monitoring Program. Data accessed from the NOAA NERRS Centralized Data Management Office website: <http://www.nerrsdata.org/>; *accessed* 12 October 2018.

NERR meteorological data and metadata can be obtained from the Research Coordinator at the individual NERR site (please see Principal Investigators and Contact Persons), from the Data Manager at the Centralized Data Management Office (please see personnel directory under the general information link on the CDMO home page) and online at the CDMO home page [www.nerrsdata.org](http://www.nerrsdata.org).  Data are available in comma delimited format.

**8) Associated researchers and projects –**

The SWMP program in place on Sapelo also includes water quality and nutrient datasets collected along with this meteorological data. Sapelo Island has a long history of maintaining research. In 1953, the University of Georgia Marine Institute (UGAMI) was formed and the island became a working laboratory for many. The research continues today with SAP NERR and UGAMI creating a unique partnership with much of the current research being done facilitated by SAP NERR and UGAMI together. Given UGAMI's long history on Sapelo, a bibliographic list of over 800 articles of current and previous research can be found on the UGAMI website: <http://www.uga.edu/ugami> and on the Sapelo Island NERR site: <http://www.sapelonerr.org> .

**II. Physical Structure Descriptors**

**9) Sensor specifications –**

Parameter: Temperature

Units: Celsius

Sensor type: Platinum resistance temperature detector (PRT)

Model #: HC2S3 Temperature and Relative Humidity Probe

Operating Temperature: -40°C to +60°C

Range: -50°C to +100°C

Accuracy: ± 0. °C with standard configuration settings

Date of Last calibration: 10/01/2015

Dates of Sensor Use: 02/02/2016 - 05/08/2018

SN 61115039

Date of Last calibration: 11/25/2015

Dates of Sensor Use: 05/08/2018 - current as of 12/31/2018

SN 0020072700

Parameter: Relative Humidity

Units: Percent

Sensor type: HC2S3 Temperature and Relative Humidity Probe

Range: 0-100% non-condensing

Accuracy at 23°C: +/- .8% RH

Temperature dependence of RH measurement: +/- 3% (-40 to 60C)

Date of Last calibration: 10/01/2015

Dates of Sensor Use: 02/02/2016 - 05/08/2018

SN 61115039

Date of Last calibration: 11/25/2015

Dates of Sensor Use: 05/08/2018 - current as of 12/31/2018

SN 0020072700

Parameter: Barometric Pressure

Units: millibars (mb)

Sensor type: Vaisala Barocap © silicon capacitive pressure sensor

Model #: CS-105

Operating Range: Pressure: 600 to 1060 mb; Temperature: -40°C to +60°C;

Humidity: non-condensing

Accuracy: ± 0.5 mb @ 20°C; +/- 2 mb @ 0°C to 40°C; +/- 4 mb @ -20°C to 45°C; +/- 6 mb @ -40°C to 60°C

Stability: ± 0.1 mb per year

Date of Last calibration: 10/14/2015

Dates of Sensor Use: 02/02/2016 - 05/08/2018

SN: J0560059

Date of Last calibration: 12/20/2016

Dates of Sensor Use: 05/08/2018 - current as of 12/31/2018

SN: M4930097

Parameter: wind speed

Units: Meters/second (m/s)

Sensor type: 12cm dia cup wheel assembly, 40 mm dia hemispherical cups

Model #: R.M. Young 03110-5 Wind Sentry

Range: 0-50m/s (112mph) gust survival; 60m/s (134mph)

Accuracy: +/-2%

Last service: 10/06/2015

Dates of Sensor Use: 02/02/2016 - 10/03/2017

SN: unknown

Last service: New in 2016

Dates of Sensor Use: 10/03/2017 - current as of 12/31/2018

SN: unknown

Parameter: wind direction

Units: degrees

Sensor type: balanced vane 16cm turning radius

Model #: R.M. Young 03110-5 Wind Sentry

Range: 360deg

Accuracy: +/- 5%

Last service: 10/06/2015

Dates of Sensor Use: 02/02/2016 - 10/03/2017

SN: unknown

Last service: New in 2016

Dates of Sensor Use: 10/03/2017 - current as of 12/31/2018

SN: unknown

Parameter: Photosynthetically Active Radiation (PAR)

LI-COR Quantum Sensor

Units: mmoles m-2 (total flux)

Sensor type: High stability silicon photovoltaic detector (blue enhanced)

Model #: LI190SB

Light spectrum waveband: 400 to 700 nm

Temperature dependence: 0.15% per °C maximum

Stability: <±2% change over 1 yr

Operating Temperature: -40°C to 65°C; Humidity: 0 to 100%

Sensitivity: typically 5 µA per 1000 µmoles s-1 m-2

Multiplier: 1.174205

Date of last calibration: 10/07/2015

Dates of Sensor Use: 02/02/2016 - 05/08/2018

SN: Q49457

Apogee Quantum Sensor

Sensor type: anodized aluminum with cast acrylic diffuser

Model #SQ110 Apogee Quantum Sensor

Light spectrum waveband: 410 to 655 nm

Temperature dependence: 0.06+/-0.06% per °C

Stability: <±2% change over 1 yr

Operating Temperature: -40°C to 70°C; Humidity 0 to 100%

Cosine Response: 45° zenith angle: +/- 2%; 75° zenith angle: +/- 5%

Sensitivity: 0.2mV per µmol s-1 m-2

Multiplier: 0.025

Last cal: 03/10/2017

Serial # SQ-110\_22481

Dates of Sensor Use: 05/08/2018 - current as of 12/31/2018

Parameter: Precipitation

Units: millimeters (mm)

Sensor type: Tipping Bucket Rain Gauge

Model #: TE525

Rainfall per tip: 0.01 inch

Operating range: Temperature: 0° to 50°C; Humidity: 0 to 100%

Accuracy: +/- 1.0% up to 1 in./hr; +0, -3% from 1 to 2 in./hr; +0, -5% from 2 to 3 in./hr

Date of Last calibration: 03/01/2017, previous calibrations: 02/02/2016, 10/08/2014

SN:50994-412

The CR1000 has 2 MB of Flash EEPROM that is used to store the Operating System. Another 128 K Flash is used to store configuration settings. A minimum of 2 MB SRAM is (4 MB optional upgrade) available for program storage (16K), operating system use, and data storage. Additional storage is available by using a compact flash card in the optional CFM100 Compact Flash Module.

SN 5203 installed from 6/26/2006 - 04/24/2015

Date CR1000 calibrated: 07/14/2014

Date CR1000 Installed: 5/5/2015 - current as of 12/31/2018

SN: 66489

**CR1000 Firmware Version (s):** unknown

Currently running CR1000 program Version:

SAPMLMET\_V6.0\_020216.CR1

SAPMLMET\_V6.0\_051018

**10) Coded variable definitions –**

[Sampling station: Sampling site code: Station code:

Marsh Landing ML sapmlmet

**11) QAQC flag definitions –**

QAQC flags provide documentation of the data and are applied to individual data points by insertion into the parameter’s associated flag column (header preceded by an F\_). During primary automated QAQC (performed by the CDMO), -5, -4, and -2 flags are applied automatically to indicate data that is above or below sensor range, or missing. All remaining data are then flagged 0, as passing initial QAQC checks. During secondary and tertiary QAQC 1, -3, and 5 flags may be used to note data as suspect, rejected due to QAQC, or corrected.

-5 Outside High Sensor Range

-4 Outside Low Sensor Range

-3 Data Rejected due to QAQC

-2 Missing Data

-1 Optional SWMP supported parameter

0 Passed Initial QAQC Checks

1 Suspect Data

2 *Open - reserved for later flag*

3 *Open - reserved for later flag*

4 Historical Data: Pre-Auto QAQC

5 Corrected Data

**12) QAQC code definitions** –

QAQC codes are used in conjunction with QAQC flags to provide further documentation of the data and are also applied by insertion into the associated flag column. There are three (3) different code categories, general, sensor, and comment. General errors document general problems with the CR1000, sensor errors are sensor specific, and comment codes are used to further document conditions or a problem with the data. Only one general or sensor error and one comment code can be applied to a particular data point, but some comment codes (marked with an \* below) can be applied to the entire record in the F\_Record column.

General Errors

GIM Instrument malfunction

GIT Instrument recording error, recovered telemetry data

GMC No instrument deployed due to maintenance/calibration

GMT Instrument maintenance

GPD Power down

GPF Power failure / Low battery

GPR Program reload

GQR Data rejected due to QA/QC checks

GSM See metadata

Sensor Errors

SDG Suspect due to sensor diagnostics

SIC Incorrect calibration constant, multiplier or offset

SIW Incorrect wiring

SMT Sensor maintenance

SNV Negative value

SOC Out of calibration

SQR Data rejected due to QAQC checks

SSD Sensor drift

SSN Not a number / unknown value

SSM Sensor malfunction

SSR Sensor removed

Comments

CAF Acceptable calibration/accuracy error of sensor

CCU Cause unknown

CDF Data appear to fit conditions

CML Snow melt from previous snowfall event

CRE\* Significant rain event

CSM\* See metadata

CVT\* Possible vandalism/tampering

CWE\* Significant weather event

**13) Other remarks/notes** –

Data are missing due to equipment or associated specific sensors not being deployed, equipment failure, time of maintenance or calibration of equipment, or repair/replacement of a sampling station platform. Any NANs in the dataset stand for “not a number” and are the result of low power, disconnected wires, or out of range readings. If additional information on missing data is needed, contact the Research Coordinator at the reserve submitting the data.

Small negative PAR values are within range of the LI-COR sensor and are due to normal errors in the sensor and the CR1000 Datalogger. The Maximum signal noise error for the LI-COR sensor is +/- 2.214 mmoles/m2 over a 15 minute interval. These values are automatically flagged and coded as <1> (CAF).

Relative Humidity data greater than 100 are within range of the sensor accuracy of +/-3% and are flagged and coded as suspect, <1> (CAF). Values greater than 103 are rejected <-3>.

Data recorded for all parameters (with the exception of cumulative precipitation) at the midnight timestamp (00:00) are the 15 minute averages and totals for the 23:45-23:59 time period of the previous day. Cumulative precipitation data at the midnight timestamp (00:00) are the sum of raw (unrounded) precipitation data from 00:00 to 23:59 of the previous day. Summing each individual 15-minute total precipitation value from the same period will result in small differences from cumulative precipitation due to rounding. It is especially important to note how data at the midnight timestamp are recorded when using January 1st and December 31st data. **Note: Cumulative precipitation is no longer available via export from the CDMO. Please contact the Reserve or the CDMO for more information or to obtain these data.**

CSM:

There were noticeable changes in PAR values following the swap to a freshly calibrated sensor (assumed to be accurate) on 05/08/2018.  The sensor (Q49457) was retired after removal and was not post calibrated to check for drift. Since drift numbers are not known all PAR data 1 year prior the sensor swap, from 05/08/2017 to 05/08/2018 11:00 are flagged and coded as <1> SSD CSM.  PAR data for the remainder of this deployment (02/02/2016 - 05/07/2017) are flagged and coded <0> CSM (unless otherwise flagged as suspect or rejected; see below for elevated nighttime flagging and coding) and users should note that drift for that period may have exceeded acceptable limits as well.  If users are comfortable assuming that drift was linear (in a real world environment it is unlikely to be entirely linear), these data may be ‘corrected’ for assumed linear drift at the user’s discretion using manufacturer’s instructions.

In addition to drift, PAR data from the beginning of the year until a sensor upgrade on 5/8 at 11:00 is also considered suspect due to consistently elevated nighttime readings. It is believed that this was caused by moisture in the sensor, times of rain or heavy dew seemed to exacerbate the problem. We are unsure how daytime readings were affected.

RH values from the beginning of the year, 01/01/2018 00:15 until 02/02/2018 00:00, when the data were flagged and coded for being out of calibration, are coded with CSM. RH values often topped off at 100%. The sensor may have started to drift prior to being considered out of calibration.

Air temperature and relative humidity are considered suspect due to being collected with an out of calibration sensor. Data are flagged as suspect, <1> SOC CSM, from 02/02/2018 00:15 - 05/08/2018 11:00.

Air temperature and relative humidity values were rejected on the following dates and times due to low minimum air temperature readings being recorded. Minimum air temperature values are not included in the dataset but are available from the Reserve. Rejected data are coded as SQR CSM.

07/11/2018 17:15

08/01/2018 02:30

08/02/2018 08:15

08/17/2018 14:30

09/09/2018 20:30

09/18/2018 21:45

10/19/2018 23:00

10/22/2018 16:45

10/25/2018 00:00

10/26/2018 16:30

12/24/2018 01:15

12/24/2018 23:45

12/29/2018 15:30

-39.3

-39.6

-39.3

-39.2

-39.2

-39.5

-39.6

-39.4

-39.6

-39.4

-39.6

-39.6

-39.5

Following maintenance to swap sensors (air temperature/relative humidity, barometric pressure, and PAR) on 05/08/2018 there was an unsuccessful attempt to upload a new CR1000 program. The station program was modified on 05/08/2018 and re-uploaded at 11:15 but an error resulted in total data rejection, <-3> CSM, until 05/10/2018 11:30.

On 05/08/2018 the LiCor PAR sensor was removed and replaced with a different model PAR sensor. Beginning with the 05/08/2018 sensor swap, SAP NERR is now using an Apogee Quantum Sensor. See the sensor specs for more detailed information. All PAR data are rejected from 05/08/2018 11:15 – 05/10/2018 11:30.

05/10/2018 11:45 – 06/30/2018 23:45 telemetered data were used to replace data that were unable to be collected from the CR1000. These data are considered suspect, <1> GIT CSM. Although only available from the Reserve, max and min air temperature and cumulative precipitation missing from the telemetry file are not available. Because data on 07/01/2018 00:00 are missing and since we are unsure of the accuracy of the 00:15 values, those data are considered suspect.